## Commands
Command | Example response | Description
!STATUS | OK > CLOSED IDLE | Returns the machine status. Drawer can be CLOSED or OPEN. Machine can be IDLE, MEASURING, and sometimes AIR CAL 405
!TEMP | OK > 0 24.8 25.0 | Returns the set temperature, an unknown temperature, and the sample temperature in that order. A set temperature of zero indicates that active temperature control is disabled.
!TEMP x1 | OK | Sets the set temperature to x1. If x1 is 0, disable active temperature control
!OPEN | OK | Opens the drawer
!CLOSE | OK | Closes the drawer
!SHAKE NOW | OK | Starts sample shaking
!SHAKE STOP | OK | Stops sample shaking
!SHAKE x1 x2 x3 x4 0 | OK | Shake plate for x1 seconds before plate read. For kinetic mode, x2 is the interval between reads in seconds, x3 is the time in seconds to wait before shaking between reads, and x4 is the duration of shaking between reads. For example, !SHAKE 5 20 16 4 0 would indicate to shake 5 seconds before the first read, wait 16 seconds, shake for 4 seconds, wait for 16 seconds, shake for 4, etc. It seems like x2 = x3+x4 in all cases, but I am not sure.
!SHAKE OFF | OK | Don't use shaking for the next sample read
!SHAKE ON | OK | Use shaking for the next sample read
!READTYPE ABSCUV | OK > readstage complete | Absorbance cuvette read type. This command can require 5-10 seconds, and is not complete until "readstage complete" is sent
!READTYPE FLUCUV | OK > readstage complete | Fluoresence cuvette read type. This command can require 5-10 seconds, and is not complete until "readstage complete" is sent
!READTYPE ABSPLA | OK > readstage complete | Absorbance plate read type. This command can require 5-10 seconds, and is not complete until "readstage complete" is sent
!READTYPE TIME x1 x2 | OK | Time resolved fluoresence read type. x1 is the start integration time in microseconds, and x2 is the stop time. Typically x1=50 and x2=850
!READTYPE FLU | OK | Fluoresence plate read type
!READTYPE LUM | OK | Luminesence plate read type
!READSTAGE TOP | OK > readstage complete | Sets the plate reader to read from the top of the well plate. This command can require 5-10 seconds, and is not complete until "readstage complete" is sent
!READSTAGE BOT | OK > readstage complete | Sets the plate reader to read from the bottom of the well plate. This command can require 5-10 seconds, and is not complete until "readstage complete" is sent
!TOP ON | OK | Unknown. This command is sometimes issued even when reading from the bottom of the plate. It might be unrelated to top and bottom reading.
!WAVELENGTH x1 x2 x3 x4 x5 x6 | OK | Sets read wavelengths for absorbance. x1 is required. x2 through x6 are optional.
!SPEED OFF | OK > normal mode | Sets "speed mode" to off. Speed mode uses stored calibration for certain fluorescence measurements. Normal mode collects calibration data right before acquisition from an internal standard, most likely.
!SPEED ON |  | Sets "speed mode" to on.
!READ | OK | Starts the measurement
!CLEAR DATA | OK | Clears the data queue
!QUEUE | OK > 1 | Returns the number of items in the transfer queue. In this example, 1.
!TRANSFER | OK > 19370123 27.2 plate L: 405 700 L: 405 4: -0.0001 -0.0001 5: -0.0000 -0.0000 L: 700 4: -0.0006 -0.0007 5: -0.0005 -0.0007 144 | Transfer one item from the queue. Format varies for each read type. First large integer is time in ticks since machine started. 5000 ticks per second.
!MODE KINETIC x1 x2 | OK | Reads x2 number of times with x1 interval in seconds
!MODE ENDPOINT | OK | Reads once immediately.
!MODE SPECTRUM x1 x2 x3 | OK | Reads spectrum data starting at wavelength x1, increment by x2, x3 number of times
!MODE EMSPECTRUM x1 x2 x3 | OK | Fixed excitation wavelegnth. Read spectrum of emission starting at x1 nm, increment x2, x3 number of times
!MODE EXSPECTRUM x1 x2 x3 | OK | Fixed emission wavelegnth. Read spectrum of excitation starting at x1 nm, increment x2, x3 number of times
!FPW x1 | OK. | Sets x1 flashes per read. 6 is typical for fluoresence. In luminesence mode, this command sets the "number of reads", with 1 minimum and 30 maximum. 30 reads corresponds to 5 seconds of PMT integration time
!AUTOPMT ON | OK | Sets PMT automatic gain to on
!AUTOPMT OFF | OK | Sets PMT automatic gain to off. Must issue another PMT command to set gain to LOW, MED, HIGH, or a numeric value. For kinetic measurement mode, use a manual PMT setting.
!PMT LOW | OK | Sets PMT gain to low
!PMT MED | OK | Sets PMT gain to medium
!PMT HIGH | OK | Sets PMT gain to high
!PMT x1 | OK | Sets PMT gain to x1 (200 to 1000 volts)
!EMWAVELENGTH x1 x2 x3 x4 x5 x6 | OK | Sets emission wavelengths
!AUTOFILTER OFF | OK | Sets automatic emission filter to off. Typically this is always done, and the software application chooses a filter setting based on the user's inputs.
!AUTOFILTER ON | OK | Sets automatic emission filter to on.
!AUTOFILTER EX OFF | OK | Sets the automatic excitation filter to off. This is only applicable for fixed emission, sweep (spectrum) excitation fluoresence
!AUTOFILTER EX ON | OK | Sets the automatic excitation filter to on. This is only applicable for fixed emission, sweep (spectrum) excitation fluoresence
!EMFILTER x1 x2 x3 x4 x5 x6 | OK | Sets emission filter 1= no filter 2=420nm 16=325nm 3=435 .. 15=695
!EXWAVELENGTH x1 x2 x3 x4 x5 x6 | OK | Sets excitation wavelengths
!WELLSCANMODE | OK > off | Returns current modal setting of well scan mode "off" or "on"
!WELLSCANMODE OFF | OK | Sets well scan mode to off
!WELLSCANMODE ON | OK | Sets well scan mode to on
!CSPEED 8 | OK | Sets carriage speed to normal
!CSPEED 1 | OK | Sets carriate speed to slow
!ORDER COLUMN | OK | Read all columns first, then go to next wavelength
!ORDER WAVELENGTH | OK | Read all wavelengths, then go to next column
!CALIBRATE ONCE | OK | Calibrate before plate read
!CALIBRATE OFF | OK | Do not run calibration. Use previously stored values.

## EmFilter
EmFilter | wavelength (nm)
1 | no filter
2 | 420
3 | 435
4 | 455
5 | 475
6 | 495
7 | 515
8 | 530
9 | 550
10 | 570
11 | 590
12 | 610
13 | 630
14 | 665
15 | 695
16 | 325

## Cuv Abs
Cuvette absorbance endpoint measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE ABSCUV |  | Absorbance cuvette
OK
>
| readstage complete | Must wait for this response, which somtimes requires 5 or more seconds
!WAVELENGTH 426 500 |  | 426nm 500nm
| OK | One wavelength is required. Up to six are allowed
!MODE ENDPOINT |  | "Endpoint" is a single measurement
OK
!SPEED OFF |  | Do not use "speed mode" (probably only applicable to plate reads in fluoresence mode)
OK
>
normal mode
!READ |  | Start the measurement
OK
!QUEUE |  | Check the queue
OK
>
1
!TRANSFER |  | Transfer the data
OK
>
13934785 25.7 cuvette
L: 426 500
L: 426 0.1114 77.3670
L: 500 0.1029 78.9028

## Cuv Kinetic
Cuvette kinetic (measurement repeated over time at specified interval)
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE ABSCUV |  | Absorbance cuvette
OK
>
| readstage complete | Must wait for this response, which somtimes requires 5 or more seconds
!WAVELENGTH 426 500 |  | 426nm and 500nm
OK
!MODE KINETIC 10 4 |  | Kinetic measurement, 10 second interval, 4 repeats
OK
!SPEED OFF |  | Speed mode off
OK
>
normal mode
!READ |  | Start the data acquisition
OK
!QUEUE |  | Check the queue
OK
>
1
!TRANSFER |  | Transfer the data
OK
>
14614364 25.8 cuvette
L: 426 500
L: 426 0.1117 77.3152
L: 500 0.1031 78.8736
!QUEUE
OK
>
1
!TRANSFER
OK
>
14664626 25.8 cuvette
L: 426 500
L: 426 0.1109 77.4577
L: 500 0.1025 78.9831

## Cuv abs spect
Cuvette absorbance spectrum measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE ABSCUV |  | Absorbance cuvette
OK
>
| readstage complete | Must wait for this response, which somtimes requires 5 or more seconds
!MODE SPECTRUM 350 10 6 |  | Start at 350nm, step 10nm, 6 steps
OK
!SPEED OFF |  | Speed mode off
OK
>
normal mode
!READ |  | Start the data acquisition
OK
!QUEUE |  | Check the queue
OK
>
1
!TRANSFER |  | Transfer the data
OK
>
16067979 25.9 cuvette
L: 350
L: 350 0.1436 71.8410
!QUEUE
OK
>
1
!TRANSFER
OK
>
16078414 25.9 cuvette
L: 360
L: 360 0.1328 73.6486

## Cuv Fluo
Cuvette fluoresence measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE FLUCUV |  | Fluoresence cuvette
OK
>
| readstage complete | Must wait for this response, which somtimes requires 5 or more seconds
!FPW 6 |  | 6 flashes per read
OK
!AUTOPMT ON |  | Automatic PMT gain
OK
!EMWAVELENGTH 525 |  | Emission wavelength 525nm
OK
!AUTOFILTER OFF |  | Automatic emission filter setting off
OK
!EMFILTER 7 |  | Emission filter manual setting 7 = 515nm.
OK
!EXWAVELENGTH 485 |  | Excitation wavelength 485nm
OK
!MODE ENDPOINT |  | "Endpoint" is a single measurement
OK
!READ |  | Start the measurement
OK
!QUEUE |  | Check the queue
OK
>
1
!TRANSFER |  | Transfer the data
OK
>
19474994 26.7 cuvette
exL: 485
emL: 525
L: 485 525
0.381

## Cuv Fluo spect Em
Cuvette Fluoresence spectrum emission
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE FLUCUV
OK
>
readstage complete
!FPW 6
OK
!AUTOPMT ON
OK
!EXWAVELENGTH 350
OK
!AUTOFILTER OFF
OK
!EMFILTER 1
OK
!MODE EMSPECTRUM 400 100 4
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
21075531 26.8 cuvette
exL: 350
emL: 400
L: 350 400
10.836
!QUEUE
OK
>
1
!TRANSFER
OK
>
21092290 26.8 cuvette
exL: 350
emL: 500
L: 350 500
6.419

## Cuv Fluo spect Ex
Cuvette Fluoresence spectrum excitation
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE FLUCUV
OK
>
readstage complete
!FPW 6
OK
!AUTOPMT ON
OK
!EMWAVELENGTH 800
OK
!AUTOFILTER OFF
OK
!EMFILTER 1
OK
!AUTOFILTER EX OFF
OK
!MODE EXSPECTRUM 400 100 4
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
22856877 26.9 cuvette
exL: 400
emL: 800
L: 400 800
336.055
!QUEUE
OK
>
1
!TRANSFER
OK
>
22886033 26.9 cuvette
exL: 500
emL: 800
L: 500 800
1.912

## Pla Abs
Plate absorption measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE ABSPLA
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!XPOS 14.380 9 12 |  | 96 well plate: Want to read block from B2 to C7. Machine must read rectangular contiguous block of wells.
| OK | 96 well plate. 14.38mm first well from edge of plate. 9mm well spacing in X direction. 12 wells
!YPOS 20.235 9 2 |  | Note: Protocol always wants to describe full "strips" (rows). So, the XPOS command is always "14.38 9 12" for 96 well plates. Use the strip command to get a subset in the X direction
| OK | 96 well plate. 20.235 (11.235 + 9) from edge to first well to read. 9mm well spacing in Y direction. 2 wells.
!SHAKE OFF
OK
!SHAKE 0 0 0 0 0
OK
!STRIP 2 6 |  | Read the strip (row) starting at well 2, and read 6 wells to the right, inclusive.
OK
!CSPEED 8 |  | Carriage speed 8 is normal
OK
!WAVELENGTH 405 500
OK
!CALIBRATE ONCE
OK
!MODE ENDPOINT
OK
!ORDER COLUMN
OK
!SPEED OFF
OK
>
normal mode
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
24615879 26.9 plate
L: 405 500
L: 405
2: -0.0000 -0.0000
3: 0.0000 -0.0000
4: 0.0000 -0.0000
5: 0.0000 -0.0000
6: 0.0000 -0.0000
7: 0.0000 -0.0001
L: 500
2: 0.0000 -0.0002
3: -0.0000 -0.0002
4: -0.0000 -0.0001
5: -0.0001 -0.0002
6: -0.0000 -0.0002
7: -0.0000 -0.0002
146

## Pla Fluo
Plate fluoresence measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE FLU
OK
!READSTAGE TOP
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!XPOS 14.380 9 12
OK
!YPOS 20.235 9 2
OK
!SHAKE OFF
OK
!SHAKE 0 0 0 0 0
OK
!STRIP 2 6
OK
!CSPEED 8
OK
!TOP ON
OK
!FPW 6
OK
!AUTOPMT ON
OK
!EMWAVELENGTH 525 600
OK
!AUTOFILTER OFF
OK
!EMFILTER 7 11
OK
!EXWAVELENGTH 485 485
OK
!PMTCAL ON
OK
!MODE ENDPOINT
OK
!ORDER COLUMN
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
27745214 27.0 plate
exL: 485 485
emL: 525 600
L: 485 525
2: 0.223 0.226
3: 0.264 0.221
4: 0.262 0.244
5: 0.273 0.245
6: 0.223 0.209
7: 0.269 0.220
L: 485 600
2: 0.218 0.187
3: 0.239 0.202
4: 0.220 0.258
5: 0.233 0.236
6: 0.215 0.206
7: 0.183 0.190
37

## Pla TRF
Plate Time Resolved Fluoresence
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE TIME 50 850 |  | Time resolved fluoresence 50 to 850 microseconds
OK
!READSTAGE TOP
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!XPOS 14.380 9 12 |  | 96 well plate: Want to read block from B3 to C7. Machine must read rectangular contiguous block of wells.
| OK | 96 well plate. 14.38mm first well from edge of plate. 9mm well spacing in X direction. 12 wells
!YPOS 20.235 9 2 |  | Note: Protocol always wants to describe full "strips" (rows). So, the XPOS command is always "14.38 9 12" for 96 well plates. Use the strip command to get a subset in the X direction
| OK | 96 well plate. 20.235 (11.235 + 9) from edge to first well to read. 9mm well spacing in Y direction. 2 wells.
!SHAKE OFF
| OK | Read the strip (row) starting at well 3, and read 5 wells to the right, inclusive.
!SHAKE 0 0 0 0 0
OK
!STRIP 2 6
OK
!CSPEED 8 |  | Carriage speed 8 is normal
OK
!TOP ON
OK
!FPW 20
OK
!AUTOPMT ON
OK
!EMWAVELENGTH 525 600
OK
!AUTOFILTER OFF
OK
!EMFILTER 7 11
OK
!EXWAVELENGTH 485 485
OK
!PMTCAL ON
OK
!MODE ENDPOINT
OK
!ORDER COLUMN
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
29941133 26.9 plate
exL: 485 485
emL: 525 600
L: 485 525
2: 505.556 488.526
3: 493.286 523.898
4: 488.542 494.231
5: 519.462 510.029
6: 496.194 494.513
7: 496.051 499.817
L: 485 600
2: 511.084 504.339
3: 518.001 500.183
4: 490.373 501.546
5: 494.754 496.492
6: 504.422 494.916
7: 499.032 520.163
181

## Pla Lum
Plate Luminosity measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE LUM
OK
!READSTAGE TOP
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!XPOS 14.380 9 12
OK
!YPOS 20.235 9 2
OK
!SHAKE OFF
OK
!SHAKE 0 0 0 0 0
OK
!STRIP 2 6
OK
!CSPEED 8
OK
!TOP ON
OK
!FPW 30
OK
!AUTOPMT ON
OK
!EMWAVELENGTH 0 578
OK
!PMTCAL ON
OK
!MODE ENDPOINT
OK
!ORDER COLUMN
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
31729551 26.9 plate
exL: f 0 f 0
emL: 0 578
L: 0 0
2: 233.578 245.843
3: 243.367 223.624
4: 241.115 239.399
5: 232.408 235.596
6: 248.986 232.973
7: 251.098 244.118
L: 0 578
2: 232.823 231.495
3: 229.757 248.876
4: 254.941 248.499
5: 222.539 232.410
6: 223.690 237.296
7: 256.785 242.938
188

## Pla Lum spect
Plate Luminosity spectrum measurement
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE LUM
OK
!READSTAGE TOP
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!XPOS 14.380 9 12
OK
!YPOS 20.235 9 1
OK
!SHAKE OFF
OK
!SHAKE 0 0 0 0 0
OK
!STRIP 2 1
OK
!CSPEED 8
OK
!TOP ON
OK
!FPW 30
OK
!AUTOPMT ON
OK
!PMTCAL ON
OK
!MODE EMSPECTRUM 300 100 5
OK
!ORDER WAVELENGTH
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
33518050 26.8 plate
exL: f 0
emL: 300
L: 0 300
2: 255.856
240
!QUEUE
OK
>
1
!TRANSFER
OK
>
33593280 26.8 plate
exL: f 0
emL: 400
L: 0 400
2: 262.319
242

## Pla Abs wellscan
Plate absorbance wellscan (read nine points in a 3x3 pattern within each well)
Command sent to device | Response | Comment
!CLEAR DATA
OK
!READTYPE ABSPLA
OK
>
readstage complete
!WELLSCANMODE
OK
>
off
!WELLSCANMODE ON
OK
!XPOS 13.247 9 12
OK
!YPOS 19.102 9 2
OK
!SHAKE OFF
OK
!SHAKE 0 0 0 0 0
OK
!STRIP 2 2
OK
!CSPEED 8
OK
!WAVELENGTH 426
OK
!CALIBRATE ONCE
OK
!MODE ENDPOINT
OK
!ORDER WAVELENGTH
OK
!SPEED OFF
OK
>
normal mode
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
42425189 26.8 plate
L: 426
L: 426
2: 0.0000 -0.0001
3: 0.0000 -0.0001
87
!XPOS 14.380 9 12
OK
!YPOS 19.102 9 2
OK
!SHAKE OFF
OK
!PMTCAL OFF
OK
!STRIP 2 2
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
42462262 26.8 plate
L: 426
L: 426
2: 0.0000 -0.0001
3: -0.0000 -0.0001
93
!XPOS 15.513 9 12
OK
!YPOS 19.102 9 2
OK
!SHAKE OFF
OK
!PMTCAL OFF
OK
!STRIP 2 2
OK
!READ
OK
!QUEUE
OK
>
1
!TRANSFER
OK
>
42498430 26.7 plate
L: 426
L: 426
2: 0.0000 -0.0000
3: 0.0000 0.0000
70
Continue reading nine positions